**SETTING UP THE ZOLL AED3®**

To ensure consistency, the Alabama Council of EMS has determined a universal configuration file for all AEDs purchased under this grant. This file will be provided to each agency to download onto a USB drive. Each agency will need to add their own Wi-Fi connection point to the AEDs. This document will guide you through importing the provided configuration file, uploading said file to the AED3, and adding your department’s Wi-Fi information.

First, let’s upload the AED3 configuration.

**Importing the AED3 Configuration Files**

* Press and hold power button for 10 seconds, then release

Screen will briefly turn white, followed by ZOLL screen

* + Select the **Settings** icon A screenshot of a computer

    Description automatically generated
  + Select the **Supervisor** menu 
  + Enter the supervisor code, then press **OK**. Your default code is **123456**
  + Insert the USB drive into the USB port, located at the rear of the device, above the battery and below the handle. Wait for USB symbol to appear at bottom, left-hand corner of screen before continuing. Should take 10-15 seconds.
  + Select **Import Files** icon 
  + Select **Configuration**, press **OK**
  + Select **Clinical Configuration**, press **OK**
  + Select **Communications Configuration**, press **OK**

The configuration import is now complete. Next, enter your agency’s Wi-Fi information. You can enter up to 25 Wi-Fi access points. For convenience, this information is available on a step-by-step video at the weblink below. Alternatively, the PowerPoint provided during training has a scannable QR code. Fast-forward to 7:15.

Link: <https://www.zoll.com/contact/elearning/videos/zoll-aed-3>

After entering and testing all needed Wi-Fi access points, this AED is not fully configured. Insert a new USB drive (or delete the previous USB drive and re-insert) to avoid using the wrong files going forward. We will now save the new configuration file (with Wi-Fi access points) to the USB.


**Exporting Your New AED3 Configuration Files With Wi-Fi Added**

* A screenshot of a computer

  Description automatically generatedReturn to the main **Supervisor** menu screen (see directions above, if needed)
  + Select **Export Files** icon 
  + Select **Configuration**, press **OK**
  + Select **Configuration**, press **Save (USB)**

You can now use the USB to configure all other devices. Follow the instructions on page one, under **Importing the AED3 Configuration File**

If you need assistance with configuring your aed3 for communications, please contact zoll technical support at (800) 348-9011.

****

